Oxygen Administration with an Oxygen Concentrator

Considerations:

- Medical oxygen is classified by the Food and Drug Administration (FDA) as a drug; therefore, a prescription from a health care provider is required.

- Portable liquid oxygen tanks can be refilled from a home-based liquid oxygen system.

- These tanks are student specific and only indicated as part of a student’s Individualized Health Care Plan (IHCP) and Emergency Action Plan (EAP)/health care provider’s order.

- An oxygen concentrator needs to be situated at least 6 inches away from walls, furniture, and drapes.

- The supplier of the equipment may be available to provide initial set-up and in-person instruction.

- Backup supply of oxygen should be available in a school setting in case of oxygen concentrator malfunction or power outage.

Supplies:

- Oxygen concentrator

- Oxygen delivery method

- Oxygen tubing

- Health care providers order

- Individualized Healthcare Plan (IHCP) and/or Emergency Action Plan (EAP)/Health care provider’s orders

- Parent/Guardian authorization form

Administering Oxygen with an Oxygen Concentrator Procedure:

- Call for assistance, if needed

- Wash hands

- Position the student to provide as much privacy as possible

- Explain procedure to student at a level they will understand

- Inspect nasal cannula and oxygen tubing to ensure it is in proper working condition

- Plug AC cord into properly grounded outlet

- Push power strip on

- In most cases an alarm will sound at start up

- Review the health care provider’s order to determine:

  - Dose of oxygen

  - Route of administration

  - Duration of administration

- Review the six rights of medication administration to ensure it is:

  - The right student

  - The right medication

  - The right dose

  - Being given at the right time

  - Being given by the right route

  - Being given for the right reason

- Ensure proper documentation of parent/guardian authorization to administer medication

- Use control knob or button to control the flow of oxygen as ordered per the health care provider’s order

- Connect the oxygen tubing to the oxygen tank

- Review the six rights of medication administration to ensure it is:

  - The right student

  - The right medication

  - The right dose

  - Being given at the right time

  - Being given by the right route

  - Being given for the right reason

- Place nasal cannula prongs into nose so they follow the curve of the nostrils

- Place tubing over and behind each ear with adjustor comfortably under the student’s chin

  - Take care not to put undue pressure on the nasal tissue from tightening the attachment too much

- Check that oxygen is flowing appropriately

- If oxygen concentrator alarms, switch student to backup oxygen source before trying to troubleshoot (see below for troubleshooting procedure)

- Follow up with the parent or guardian and health care provider, as needed

- Document medication administration in the student’s medication administration log

Procedure for Checking and Cleaning Oxygen Concentrator and Tubing:

- Oxygen concentrator should be cleaned at least once per week

- Wipe down the oxygen concentrator with a clean damp cloth

- Check student’s health care orders or discuss with health care provider regarding how often oxygen tubing is replaced

- Check, clean, and change out the filters, as instructed by health care provider team and oxygen concentrator manufacturer’s manual

Troubleshooting oxygen concentrator alarms:

- Be sure to switch student to backup oxygen source before troubleshooting

- Check to make sure the power source is plugged into a working outlet

- Check air intake area to make sure that it is not blocked

- Make sure filters are cleaned and placed properly

- Check to make sure the flow control is not set at zero or above the highest setting

- Check to make sure the oxygen tubing is not kinked or blocked

- If none of these steps fix the issue, check the oxygen concentrator’s manufacturer’s manual

- If all else fails:

  - Turn the concentrator off

  - Wait 10 minutes for it to cool down

  - Turn concentrator back on

  - If still unable to get concentrator to work properly call parents/guardians

References:

Lynn, P. (2019). Administering oxygen by nasal cannula. In Skill checklists for Taylor’s clinical nursing skills. A nursing process approach. (5th ed.). (p. 312).

Pediatric Home Services. (2007). Deliver oxygen with an oxygen tank (Cómo administrar oxígeno con un tanque de oxígeno). [video]. Retrieved June 6, 2023, from https://www.pediatrichomeservice.com/tips-how-tos/deliver-oxygen-with-an-oxygen-tank/?play=1

Pediatric Home Services. (2007). Oxygen concentrator (Concentrador de oxígeno). [video]. Retrieved June 6, 2023, from https://www.pediatrichomeservice.com/tips-how-tos/deliver-oxygen-with-an-oxygen-tank/?play=1

Pediatric Home Services. (2007). Oxygen connections (Conexiones de oxígeno) [video]. Retrieved June 6, 2023, from https://www.pediatrichomeservice.com/tips-how-tos/deliver-oxygen-with-an-oxygen-tank/?play=1

Perry, A. G., Potter, P.A., Ostendorf, W., & Laplante, N. (2021). Applying an oxygen-delivery device. In Clinical nursing skills and techniques. (10th ed.). (pp. 712-719). St. Louis, MO: Elsevier.

Perry, A. G., Potter, P.A., Ostendorf, W., & Laplante, N. (2021). Using home oxygen equipment. In Clinical nursing skills and techniques. (10th ed.). (pp. 1248-1255). St. Louis, MO: Elsevier.